# End-to-End Evaluation of RF-DETR-Based

# Microscopy Specimen Quality Assessment

Object detection • independent threshold calibration • final expert-referenced downstream evaluation

Generated 29 July 2026

| Item | Locked value |
| --- | --- |
| Detector | RFDETR2XLarge, HPO configuration 009, EMA checkpoint |
| Internal model resolution | 880 × 880 pixels |
| SAHI full-FOV slice size | 640 × 640 pixels with 20% overlap |
| Class confidence thresholds | Leucocyte=0.36; Squamous Epithelial Cell=0.35 |
| Final downstream set | 98 FOVs from 34 samples |
| Code provenance | 27c95e88 — https://github.com/Mmose95/AIPoweredMicroscope/commit/27c95e8875da3f11b486a03f522a0a27e82ab529 |

## Headline outcomes

| Evaluation stage | Primary result |
| --- | --- |
| Held-out object detection | COCO AP@50:95 0.6179; AP@50 0.8685 |
| Independent calibration | Macro F1 0.9335 at Leucocyte=0.36 and Epithelial=0.35 |
| Final downstream task | Accuracy 58.16%; balanced accuracy 57.11%; macro F1 0.5678 |

# Contents

Headline outcomes	1

Contents	2

1. Executive summary	3

2. Evaluation design and data separation	3

3. Final model and locked inference settings	3

3.1 Model selection and training	3

3.2 Locked downstream inference	4

3.3 Count-to-quality rule	4

4. Held-out object-detection evaluation	4

Per-class standard COCO metrics	5

Locked operating point at IoU=0.50	5

5. Independent confidence-threshold calibration	7

6. Final downstream quality evaluation	9

6.1 Cohort flow and runtime	9

6.2 Overall downstream performance	10

6.3 Class-level performance	10

6.4 Error direction and detected-count patterns	11

7. Integrated interpretation	12

8. Limitations	12

9. Reproducibility and result artifacts	12

Appendix A. Downstream exclusions	14

Appendix B. Misclassified downstream FOVs	14

# 1. Executive summary

This report consolidates the complete three-stage evaluation of the final two-class RF-DETR model. The detector was first assessed on a held-out object-detection test split. Class-specific confidence thresholds were then selected on a separate calibration set. Those thresholds and all downstream inference settings were locked before the expert-referenced downstream FOV labels were analyzed.

The detector achieved COCO AP@50:95 of 0.6179 and AP@50 of 0.8685 on 331 held-out 640 × 640 patches containing 4,004 objects.

Independent calibration on 53 patches containing 776 objects selected confidence thresholds of 0.36 for Leucocyte and 0.35 for Squamous Epithelial Cell; joint macro F1 was 0.9335.

The final downstream evaluation included 98 expert-labeled FOVs after four prespecified exclusions. Accuracy was 58.16%, balanced accuracy was 57.11%, and macro F1 was 0.5678.

Downstream errors were predominantly conservative: 33 of 41 errors assigned a worse quality category than the clinical reference, while 8 assigned a more favorable category.

No threshold or quality-rule parameter was changed after the downstream labels were inspected.

# 2. Evaluation design and data separation

The study design deliberately separates model fitting, standard detector testing, threshold calibration, and final downstream testing. This prevents downstream test performance from influencing the confidence thresholds used to decide whether detections enter the count-based quality rule.

| Partition/stage | Samples | Images | Leucocyte | Epithelial | Purpose |
| --- | --- | --- | --- | --- | --- |
| Detector training | 28 | 829 | 10978 | 1890 | Model fitting |
| Detector validation | 9 | 291 | 3430 | 639 | HPO/model selection |
| Detector test | 11 | 331 | 2851 | 1153 | Held-out object detection |
| Calibration | 11 | 53 | 572 | 204 | Threshold selection only |
| Downstream test | 34 | 98 | — | — | Final FOV quality evaluation |

The detector split was sample-wise (28/9/11 samples for train/validation/test). The downstream workbook contained 102 non-header image rows: 98 were eligible, one image was marked as not present in the annotation project, and three were considered too difficult for a reliable clinical classification.

# 3. Final model and locked inference settings

## 3.1 Model selection and training

| Parameter | Value |
| --- | --- |
| Selected configuration | HPO configuration 009 |
| Architecture | RFDETR2XLarge |
| Checkpoint | checkpoint_best_ema.pth |
| Internal resolution | 880 |
| Input data mode | 640 pixel source patches |
| Queries | 600 |
| Batch size / gradient accumulation | 6 / 1 |
| Learning rate | 8.0e-05 |
| Weight decay / dropout | 0.0007 / 0.15 |
| Multi-scale training | True |
| Maximum epochs / best epoch | 160 / 8 |
| Validation AP@50 | 0.9086 |
| Validation mAP@50:95 | 0.7364 |
| Random seed | 42 |
| Patch size / attention windows | 20 / 2 |

The microscopy patches remained 640 × 640 at the data interface. RF-DETR internally resized them to the selected 880 × 880 model resolution. The same distinction was retained for downstream SAHI: the slice is 640 × 640, while the model internally operates at 880.

## 3.2 Locked downstream inference

| Setting | Locked value |
| --- | --- |
| SAHI slice | 640 × 640 |
| Slice overlap | 20% horizontal and vertical |
| Standard whole-image prediction | False |
| Class confidence thresholds | Leucocyte=0.36; Squamous Epithelial Cell=0.35 |
| Within-class postprocess | GREEDYNMM, IOU=0.50, class-aware |
| Cross-class duplicate suppression | IOS≥0.75, IoU≥0.35, area ratio≥0.55 |

## 3.3 Count-to-quality rule

| Component | Observed count | Score |
| --- | --- | --- |
| Leucocyte | <10 | −1 |
| Leucocyte | 10–25 | 0 |
| Leucocyte | 26–50 | +1 |
| Leucocyte | >50 | +2 |
| Squamous epithelial | <10 | 0 |
| Squamous epithelial | 10–25 | −1 |
| Squamous epithelial | >25 | −2 |

The component scores were summed. An FOV was Qualified when the total score was at least +1 and fewer than 10 epithelial cells were detected; it was Not Qualified when the total was −1 or lower; all remaining cases were Partially Qualified.

# 4. Held-out object-detection evaluation

Standard COCO ranking metrics were calculated from detections retained at the numerical score floor of 0.001. They do not use the calibrated thresholds. The confusion matrix and operating-point precision/recall/F1 are threshold-dependent and use the independently calibrated class-specific thresholds at IoU=0.50.

| Metric | Overall result |
| --- | --- |
| COCO AP@50:95 | 0.6179 |
| AP@50 | 0.8685 |
| AP@75 | 0.7048 |
| AR@1 | 0.1082 |
| AR@10 | 0.4598 |
| AR@100 | 0.7249 |

## Per-class standard COCO metrics

| Class | AP@50:95 | AP@50 | AP@75 | AR@100 |
| --- | --- | --- | --- | --- |
| Leucocyte | 0.5094 | 0.7950 | 0.5665 | 0.6499 |
| Squamous Epithelial Cell | 0.7265 | 0.9421 | 0.8430 | 0.7999 |

## Locked operating point at IoU=0.50

| Thresholds | TP | FP | FN | Precision | Recall | F1 | FP/image |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Leucocyte=0.36; Squamous Epithelial Cell=0.35 | 3345 | 891 | 659 | 0.7897 | 0.8354 | 0.8119 | 2.692 |

Figure 1. Object-detection confusion matrix at the locked class-specific confidence thresholds and IoU=0.50. Rows are ground truth; columns are predictions. Background indicates misses or unmatched detections.

Figure 2. Detector AP across IoU thresholds. The standard headline COCO result remains AP@50:95.

The test set contained 2,851 annotated leucocytes and 1,153 squamous epithelial cells. At the locked operating point, the confusion matrix recorded 2,296 correctly localized leucocytes and 1,046 correctly localized epithelial cells. Epithelial cells showed higher localization performance than leucocytes across AP@50:95, AP@50, AP@75, and AR@100.

The saved test threshold sweep is descriptive sensitivity analysis only. Its test-set maximum was not used to select or revise the operating thresholds and is intentionally not reported as a selected operating point.

# 5. Independent confidence-threshold calibration

Calibration used 53 images from 11 tasks, with 572 leucocyte and 204 squamous epithelial annotations. Predictions were matched at IoU=0.50. Thresholds from 0.00 to 0.95 were evaluated at 0.01 increments, with extra observed-score candidates, and macro F1 was the prespecified selection criterion.

| Class | Threshold | TP | FP | FN | Precision | Recall | F1 | Jaccard |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Leucocyte | 0.36 | 513 | 57 | 59 | 0.9000 | 0.8969 | 0.8984 | 0.8156 |
| Squamous Epithelial Cell | 0.35 | 201 | 10 | 3 | 0.9526 | 0.9853 | 0.9687 | 0.9393 |

| Joint metric | Result |
| --- | --- |
| TP / FP / FN | 714 / 67 / 62 |
| Micro precision | 0.9142 |
| Micro recall | 0.9201 |
| Micro F1 | 0.9171 |
| Macro precision | 0.9263 |
| Macro recall | 0.9411 |
| Macro F1 | 0.9335 |
| Macro Jaccard | 0.8774 |

Figure 3. Leucocyte precision, recall, F1, and Jaccard across confidence thresholds. The selected F1-maximizing threshold was 0.36.

Figure 4. Squamous epithelial cell calibration curves. The selected F1-maximizing threshold was 0.35.

Figure 5. Joint macro-F1 surface for the two class-specific confidence thresholds.

The calibrated thresholds were subsequently treated as fixed method parameters. They were applied to the threshold-dependent detector results and carried unchanged into downstream SAHI inference.

# 6. Final downstream quality evaluation

## 6.1 Cohort flow and runtime

| Item | Count/result |
| --- | --- |
| Workbook rows reviewed | 102 |
| Eligible FOVs analyzed | 98 |
| Represented samples | 34 |
| Excluded FOVs | 4 |
| Qualified references | 42 |
| Partially Qualified references | 26 |
| Not Qualified references | 30 |
| Total evaluation time | 125.3 s |
| Mean detector time per FOV | 1.14 s |

All 98 eligible images completed without inference failure, and an annotated overlay was produced for every included FOV.

## 6.2 Overall downstream performance

Uncertainty intervals are percentile 95% confidence intervals from 10,000 cluster-bootstrap replicates, resampling the 34 source samples with replacement (seed 42). This preserves within-sample FOV correlation better than an image-level bootstrap.

| Metric | Estimate | Cluster-bootstrap 95% CI |
| --- | --- | --- |
| Accuracy | 0.5816 | 0.4343–0.7253 |
| Balanced accuracy | 0.5711 | 0.4357–0.7085 |
| Macro F1 | 0.5678 | 0.4232–0.6972 |
| Cohen's kappa | 0.3796 | 0.1756–0.5752 |
| Quadratic-weighted kappa | 0.5779 | 0.3583–0.7669 |

| Additional metric | Result |
| --- | --- |
| Weighted precision | 0.6730 |
| Weighted recall | 0.5816 |
| Weighted F1 | 0.5977 |
| Mean absolute class error | 0.4898 |
| Within-one-class accuracy | 0.9286 |

## 6.3 Class-level performance

| Class | Support | TP | FP | FN | Precision | Recall | F1 | Specificity |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Qualified | 42 | 25 | 0 | 17 | 1.0000 | 0.5952 | 0.7463 | 1.0000 |
| Partially Qualified | 26 | 10 | 18 | 16 | 0.3571 | 0.3846 | 0.3704 | 0.7500 |
| Not Qualified | 30 | 22 | 23 | 8 | 0.4889 | 0.7333 | 0.5867 | 0.6618 |

Figure 6. Final downstream confusion matrix. Rows are clinical reference labels and columns are locked pipeline predictions.

The pipeline correctly classified 57 of 98 FOVs. Qualified predictions had precision 1.00: no clinically Partially Qualified or Not Qualified image was predicted as Qualified. However, Qualified recall was 0.595, reflecting a conservative acceptance threshold. Partially Qualified was the most difficult class (F1 0.370), and Not Qualified recall was 0.733.

## 6.4 Error direction and detected-count patterns

| Outcome | FOVs | Percentage of all FOVs |
| --- | --- | --- |
| Exact class agreement | 57 | 58.2% |
| Predicted more severe than reference | 33 | 33.7% |
| Predicted more favorable than reference | 8 | 8.2% |
| Two-category errors | 7 | 7.1% |

| Clinical class | n | Leucocyte median (IQR) | Epithelial median (IQR) |
| --- | --- | --- | --- |
| Qualified | 42 | 102.5 (17.5–297.2) | 2.5 (2.0–7.0) |
| Partially Qualified | 26 | 11.0 (4.2–90.8) | 19.0 (9.2–27.2) |
| Not Qualified | 30 | 19.5 (7.2–43.0) | 36.5 (26.0–48.8) |

Across all FOVs, 56,482 raw sliced predictions were reduced to 12,220 candidates by the locked confidence thresholds and within-class SAHI postprocess. Cross-class suppression removed 31 additional duplicates, leaving 12,189 detections (10,112 leucocytes and 2,077 epithelial cells).

# 7. Integrated interpretation

The final RFDETR2XLarge detector showed strong held-out localization performance, especially for squamous epithelial cells. Leucocyte detection remained the weaker component.

Independent calibration produced high object-level F1 and provided a defensible, prespecified operating point for any threshold-dependent result.

The lower downstream performance does not invalidate the detector or calibration results. It measures the entire chain: sliced full-FOV inference, duplicate handling, class-specific thresholding, count aggregation, and the deterministic clinical quality rule.

Most downstream errors were conservative. This explains why Qualified precision was perfect while Qualified recall was substantially lower.

The Partially Qualified category showed the greatest overlap with the deterministic rule boundaries and was most frequently assigned to Not Qualified.

The final downstream set must remain untouched for optimization. Any future revision of thresholds or the count-to-quality rule should be developed on a new training/development cohort and then tested on an independent external set.

# 8. Limitations

The downstream reference is the final expert-review workbook. Details of the review/consensus procedure are not fully encoded on every workbook row and should be described from the clinical study records in the manuscript.

Four of 102 reviewed rows were excluded: one was not present in the annotation project and three were too difficult for a reliable clinical classification.

The downstream analysis contains multiple FOVs from some source samples. Cluster-bootstrap intervals account for this grouping, but the number of represented samples remains 34.

Confidence thresholds were calibrated on 640 × 640 object-level patches, whereas the downstream task uses overlapping SAHI slices extracted from full FOVs. The locked approach preserves test independence but does not remove this deployment-domain difference.

The count-to-quality rule is deterministic and may be a larger source of downstream disagreement than residual detector error. This report does not tune that rule on the final test labels.

Confidence intervals quantify resampling uncertainty in this dataset; they do not establish external generalizability.

# 9. Reproducibility and result artifacts

| Artifact | Location |
| --- | --- |
| Code commit | https://github.com/Mmose95/AIPoweredMicroscope/commit/27c95e8875da3f11b486a03f522a0a27e82ab529 |
| Object-detection summary | C:\Users\SH37YE\Desktop\PhD_Code_github\AIPoweredMicroscope\EvaluationOutput\TwoClass_HPO009_EMA_TestEval_20260701-140821\eval_summary.json |
| Object predictions | C:\Users\SH37YE\Desktop\PhD_Code_github\AIPoweredMicroscope\EvaluationOutput\TwoClass_HPO009_EMA_TestEval_20260701-140821\predictions_coco.json |
| Calibration summary | C:\Users\SH37YE\Desktop\PhD_Code_github\AIPoweredMicroscope\EvaluationOutput\RFDETR_ThresholdCalibration_20260701-140419\threshold_calibration_summary.json |
| Calibration predictions | C:\Users\SH37YE\Desktop\PhD_Code_github\AIPoweredMicroscope\EvaluationOutput\RFDETR_ThresholdCalibration_20260701-140419\calibration_predictions.json |
| Downstream summary | C:\Users\SH37YE\Desktop\PhD_Code_github\AIPoweredMicroscope\DownstreamOutput_Downstream\20260729-123001\downstream_evaluation_summary.json |
| Per-FOV downstream predictions | C:\Users\SH37YE\Desktop\PhD_Code_github\AIPoweredMicroscope\DownstreamOutput_Downstream\20260729-123001\downstream_predictions.csv |
| Downstream exclusions | C:\Users\SH37YE\Desktop\PhD_Code_github\AIPoweredMicroscope\DownstreamOutput_Downstream\20260729-123001\downstream_exclusions.csv |
| Downstream overlays | C:\Users\SH37YE\Desktop\PhD_Code_github\AIPoweredMicroscope\DownstreamOutput_Downstream\20260729-123001\overlays |

The report was generated directly from the saved machine-readable artifacts above. All headline values can therefore be traced to a JSON or CSV result file.

# Appendix A. Downstream exclusions

| Workbook row | Image | Reason | Comment |
| --- | --- | --- | --- |
| 77 | Sample91 - BF.3_10.tif | not_in_cvat_project | — |
| 88 | Sample95 - BF.15_5.tif | too_difficult_for_clinical_review | for svær at vurdere uden at antage for meget |
| 89 | Sample95 - BF.5_5.tif | too_difficult_for_clinical_review | for svær at vurdere uden at antage for meget |
| 90 | Sample95 - BF.9_10.tif | too_difficult_for_clinical_review | for svær at vurdere uden at antage for meget |

# Appendix B. Misclassified downstream FOVs

The table lists all downstream disagreements. Complete results for all 98 FOVs, including correct classifications and inference diagnostics, are retained in downstream_predictions.csv.

| Image | Clinical | Predicted | Leu | Epi | Rule score |
| --- | --- | --- | --- | --- | --- |
| Sample100 - BF.10_10.tif | Not Qualified | Partially Qualified | 176 | 45 | 0 |
| Sample100 - BF.15_9.tif | Not Qualified | Partially Qualified | 237 | 26 | 0 |
| Sample100 - BF.4_9.tif | Not Qualified | Partially Qualified | 30 | 10 | 0 |
| Sample100 - BF.8_8.tif | Not Qualified | Partially Qualified | 239 | 39 | 0 |
| Sample102 - BF.12_4.tif | Partially Qualified | Not Qualified | 7 | 7 | -1 |
| Sample104 - BF.12_0.tif | Not Qualified | Partially Qualified | 54 | 89 | 0 |
| Sample105 - BF.2_12.tif | Not Qualified | Partially Qualified | 11 | 1 | 0 |
| Sample106 - BF.11_1.tif | Qualified | Not Qualified | 5 | 8 | -1 |
| Sample106 - BF.12_0.tif | Qualified | Not Qualified | 3 | 2 | -1 |
| Sample106 - BF.19_3.tif | Partially Qualified | Not Qualified | 5 | 7 | -1 |
| Sample106 - BF.7_2.tif | Partially Qualified | Not Qualified | 21 | 15 | -1 |
| Sample106 - BF.9_1.tif | Qualified | Partially Qualified | 16 | 3 | 0 |
| Sample77 - BF.10_14.tif | Not Qualified | Partially Qualified | 217 | 50 | 0 |
| Sample81 - BF.4_12.tif | Not Qualified | Partially Qualified | 286 | 60 | 0 |
| Sample83 - BF.14_7.tif | Partially Qualified | Not Qualified | 4 | 8 | -1 |
| Sample83 - BF.15_12.tif | Partially Qualified | Not Qualified | 0 | 5 | -1 |
| Sample83 - BF.16_8.tif | Partially Qualified | Not Qualified | 3 | 5 | -1 |
| Sample83 - BF.7_7.tif | Qualified | Partially Qualified | 11 | 8 | 0 |
| Sample85 - BF.10_8.tif | Partially Qualified | Not Qualified | 3 | 20 | -2 |
| Sample85 - BF.11_10.tif | Partially Qualified | Not Qualified | 10 | 13 | -1 |
| Sample87 - BF.9_4.tif | Qualified | Not Qualified | 3 | 2 | -1 |
| Sample89 - BF.16_6.tif | Qualified | Partially Qualified | 365 | 11 | 1 |
| Sample90 - BF.4_9.tif | Qualified | Partially Qualified | 388 | 23 | 1 |
| Sample91 - BF.15_0.tif | Partially Qualified | Not Qualified | 8 | 35 | -3 |
| Sample91 - BF.17_7.tif | Qualified | Partially Qualified | 14 | 7 | 0 |
| Sample91 - BF.3_8.tif | Partially Qualified | Not Qualified | 18 | 69 | -2 |
| Sample91 - BF.3_9.tif | Partially Qualified | Not Qualified | 4 | 65 | -3 |
| Sample92 - BF.11_0.tif | Qualified | Not Qualified | 5 | 0 | -1 |
| Sample92 - BF.5_12.tif | Qualified | Partially Qualified | 99 | 11 | 1 |
| Sample92 - BF.7_4.tif | Qualified | Partially Qualified | 12 | 2 | 0 |
| Sample93 - BF.14_8.tif | Partially Qualified | Not Qualified | 1 | 21 | -2 |
| Sample93 - BF.9_4.tif | Partially Qualified | Not Qualified | 0 | 5 | -1 |
| Sample96 - BF.14_10.tif | Qualified | Partially Qualified | 22 | 0 | 0 |
| Sample97 - BF.11_11.tif | Qualified | Partially Qualified | 22 | 7 | 0 |
| Sample97 - BF.12_11.tif | Partially Qualified | Not Qualified | 7 | 16 | -2 |
| Sample98 - BF.4_3.tif | Partially Qualified | Not Qualified | 16 | 18 | -1 |
| Sample98 - BF.8_9.tif | Partially Qualified | Not Qualified | 7 | 14 | -2 |
| Sample99 - BF.1_5.tif | Qualified | Partially Qualified | 24 | 3 | 0 |
| Sample99 - BF.12_8.tif | Qualified | Not Qualified | 9 | 2 | -1 |
| Sample99 - BF.14_11.tif | Qualified | Not Qualified | 8 | 6 | -1 |
| Sample99 - BF.6_12.tif | Qualified | Not Qualified | 1 | 1 | -1 |